## Slide 1: RICH Control Manual
January 22, 2024

## Slide 2
RICH front-end is sitting in the beam of the secondary particles. The radiation damage causes the malfunction of FE . Dead tiles appear. The damage is not permanent. It can be recovered by switching LV OFF/ON or reloading the RICH front-end FPGA
We don’t need to perform the RICH full recovery daily any more. DAQ is reloading FPGA memory during prestart automatically
In case you got RICH alarm do the following:
  Continue data taking if DAQ is working.
  Start new run with CANCEL-RESET
  With high probability the problem will be fixed automatically
  If the problem is not solved do RICH full recovery.
  If it does not not help call RICH expert

## Slide 3: RICH mainframe remote reboot
In case of communication lost with the RICH mainframe try to reboot it remotely
To reboot only the CPU: caenhvReset.py --soft hvrich1 (sector 4)
caenhvReset.py --soft hvrich2 (sector 1)
To power cycle the whole thing, causing all voltages to go to zero: caenhvReset.py --hard hvrich1 (sector 4)
caenhvReset.py --hard hvrich2 (sector 1)

## Slide 4
Press RICH on clascss menu
Chose RICH Overview
RICH GUI overview

## Slide 5: RICH GUI
Voltage
Control RICH HV and LV
Temperature Map
Shows the temperature of the RICH electronic boards
Scaler Map
Presents the rate of the MAPMT pixels
Hard Interlock
Control the RICH interlock
Soft Interlock
Control the max temperature of the FPGA chips

## Slide 6: Soft Interlock
Max temperature has to be less than 75 C
Reset the interlock if necessary

## Slide 7: Temperature Map
Max temperature has
to be less than 75 C
Soft interlock switches off
The RICH HV and LV if t>75C0
All tiles have to be present
except Tile 21 in sector 1

## Slide 8: RICH scalers
The plot presents the
average rate of the MAPMT
pixels
ALL MAPMTs have to
be present except Tile 21
in sector 1

## Slide 9
Hard Interlock
controls the
temperature
and humidity
inside the
RICH detector
Press this button
to view the sector 1
panel
Press this button
to view the sector 4
panel

## Slide 10: Electronic Volume Interlock
Check the temperature and humidity

## Slide 11: Nitrogen Volume Interlock
Check the temperature and humidity